## Slide 1: BIOLOGICAL VARIATION FOR N-TERMINAL PRO-BNP
Jones GRD1, Bennett A2,3, Boscato L1, Macdonald P2, Brien J3.
Departments of Chemical Pathology1, Cardiology2 and Pharmacy3, St Vincent’s Hospital, Sydney, Australia.

## Slide 2: Background
N-terminal proBNP (proBNP) is a plasma biomarker approved for use in the diagnosis and monitoring of patients with congestive cardiac failure.
Within-individual variation (CVwi) is an important parameter for assessing required analytical precision and calculating critical changes.
This data is currently available for proBNP only in healthy individuals1.
  Average CVwi = 36%
In this study we evaluate the within-individual biological variation of proBNP in patients with stable heart failure.

## Slide 3: Methods 1: Patients
Patients with symptomatic heart failure were identified during hospital admission.
Samples for proBNP were collected 2 weeks, 3 months and 6 months after discharge.
Patients were considered to have stable heart failure if, during 12 months from discharge, they:
  1. Remained alive.
  2. Did not undergo cardiac transplant.
  3. Were not re-admitted with a diagnosis of heart failure.
  4. Did not change their number of cardiac medications in 6 months.
A looser definition of stable heart failure was assessed
  Criteria 1 - 3, but allowing a single change in the number of cardiac medications during 6 months.

## Slide 4: Methods 2: Measurements
Samples were collected into K2EDTA tubes, centrifuged and plasma stored at -20 degrees until analysis.
proBNP was measured on an Elecsys 2010 analyser
Within-run CV was < 2% and total CV< 6% at all levels tested2.
All samples for a single patient were analysed in the same analytical run.

## Slide 5: Patient Demographics
After initial exclusions, 17 patients were considered to have stable heart failure (out of a total of 102).
Average age of included patients: 76 (range 62-91)
The median of all proBNP results was 2280 ng/L
  (range: 380 to 10,440 ng/L)
These results are consistent with the clinical diagnoses of NHYA class 2 or 3 heart failure.
CVs were calculated from the 3 data points for each patient.
Data from 2 patients was excluded as statistical outliers. (CV>1 SD from next result)
  One patient known not to be compliant with medication.
  One patient with a single result 7 x other two results.

## Slide 6: Within-Person Variation
Figure. Changes in proBNP over time in 17 patients with stable heart failure. The two patients with data removed due to high CVs are shown in red.
proBNP (ng/L)

## Slide 7: Results
Stable Heart Failure (2 outliers excluded) - 15 patients:
  Average proBNP was 2280 pg/mL
  Average within-person CV was 23%, range: 3 – 44%
  (excluded CVs were 66% and 119%)
Stable Heart Failure (no outlier exclusion) - 17 patients:
  Average proBNP was 3074 pg/mL
  Average within-person CV was 32% (range: 3 – 119%)
Stable Heart Failure (1 medication change) - 34 patients:
  Average proBNP result 2447 pg/mL
  The average within-person CV was 36% (range 3 - 119%)

## Slide 8: Interpretation
With the stated between-assay imprecision of 6% and the critical difference at the 95% level for each estimate of CVwi are as follows:
  CVwi of 24% - CD is a 68% change in concentration.
  CVwi of 36% - CD is 100% change in concentration.
Using the Roche assay a doubling or halving of proBNP results in a patient with heart failure is highly likely to indicate a significant change in proBNP concentration..
The Roche proBNP assay is performing better than the optimal precision level
  Optimal total precision < 0.25 x CVwi.3
  Elecsys total precision < 6%; target: 6 - 9%

## Slide 9: Discussion
Within-person biological variation has been extensively studied for many analytes in healthy people.
Fewer studies have been performed on people with disease.
A major difficulty is the definition of stable disease.
We have obtained a range of results for average CVwi depending on the definition of “stable disease”.
Our range of CVwi overlaps that reported in healthy individuals1.
The small number of samples per patient may account for the wide ranges seen.
Log transformation of the data allows better parametric description, but may be less useful in clinical practice.

## Slide 10: Conclusions
Using the tightest definition of stable heart failure the average within-individual biological variation has a CV of 24%, although higher values may be supported by the data.
Depending on definition of stable heart failure, the critical change value is between 70% and 100%.
The Roche assay clearly meets precision targets based on biological variation.
REFERENCES
1. Wu A et al. Am J Cardiol. 2004
2. Jones GRD, Boscato L, Schneider H-G, Clin Biochem Rev 2002 (abstract).
3. Fraser CG et al. Ann Clin Biochem 1997;34:8-12.
ACKNOWLEDGEMENT
We acknowledge Roche Diagnostics Australia for reagents and equipment.