## Slide 1: LABORATORY USE OF THE ABBOTT MEDISENSE METER FOR BETA-HYDROXYBUTYRATE MEASUREMENT.
GRD Jones, A Screnci, P Graham
Chemical Pathology, St Vincent’s Hospital, Darlinghurst, NSW 2010.

## Slide 2: Background
Keto-acidosis is a potentially life-threatening complication of Type 1 diabetes mellitus.
Beta-Hydroxybutyrate (BOHB) is the best marker for this complication1.
Current tests for ketones are wet chemistry laboratory tests for BOHB or dipstick tests for acetoacetate.
Low volume testing on automated analysers is not cost effective.
Abbott has developed a near-patient test for BOHB on the Medisense meter.
We evaluate the Abbott test for use in the laboratory.

## Slide 3: Requirements for Laboratory Testing
The Medisense meter is designed for near-patient testing using a whole blood fingerprick sample.
In the laboratory samples are subject to other variables:
  Effects of different sample anticoagulants
  Delays in sample arrival
  Requests for add-on tests on old samples
There is also the opportunity to improve control of such factors as sample volume and to perform additional functions such as sample dilution.

## Slide 4: Aims
To evaluate the Abbott Medisense meter for BOHB measurement in a laboratory environment.
This includes assessment of:
  Sample types
  Precision
  Sample stability
  Effect of short sample volume
  Sample Dilution
Note: Accuracy has previously been addressed in the literature1,2 and does not form part of this evaluation.

## Slide 5: Medisense BOHB Assay
Product Insert Claims
  Immediate measurement of whole blood
  Heparin whole blood suitable and stable for 30 mins
  Serum or plasma is not recommended.
  Fluoride oxalate whole blood unsuitable.
  Other sample types not mentioned
  Suitable for haematocrit range 30% to 60%.
  5 uL sample volume
  Range 0.0 - 6.0 mmol/L
  Results in 30 seconds

## Slide 6: Methods
Sample Types
  The following sample types were evaluated:
  EDTA whole blood (WB), heparin WB, EDTA plasma, heparin plasma, Fluoride oxalate WB.
Stability
  Sample stability was assessed room temperature and 4°C.
Dilution
  Linearity of whole blood dilution was assessed.
Sample volume
  Accuracy at low sample volume additions.
Acceptance Criteria for this study
  Up to +/- 0.7 mmol/L below 3.5 mmol/L, then 20%
  This criteria allows clear separation of DKA from normals3

## Slide 7: Sample Comparison
Heparin Whole Blood BOHB (mmol/L)
BOHB (mmol/L)
BOHB Difference (mmol/L)
Figure 2. Scatter plot (A) and Difference plot (B) for EDTA WB (blue diamonds, ) and Fluoride Oxalate WB (red squares, ) compared to heparin WB. The acceptance limits (+/- 0.7 up to 3.5 mmol/L, then 20%) are shown in purple. Solid line in A is the line of identity, dashed line is the line of best fit for EDTA samples.
Interpretation: EDTA and Fluoride Oxalate WB results are equivalent to Heparin WB results.
A
B

## Slide 8: Precision
Figure 3. Within-run precision was evaluated by duplicate testing of whole blood samples. The graph shows the difference between paired paired measurements of EDTA WB (blue diamonds, ) and heparin WB (green squares, ). The overall CV was 8% for EDTA and 7% for heparin samples (highest EDTA outlier was excluded from CV calculation).
Interpretation: Whole blood precision is acceptable.
BOHB (mmol/L)
Difference (mmol/L)

## Slide 9: Sample Stability
Figure 4. Sample stability at Room temperature (A) and 4 degrees (B) for EDTA WB (blue diamonds, ), heparin WB (green triangles, ) and Fluoride oxalate WB (red squares, ). Change is expressed as a fraction of the Acceptance Criteria (AC) on the Y axis.
Interpretation: Most samples are stable for at least 24 hours at room temperature and 3 days at 4 degrees.
A. room temperature
B. 4°C
Time (hours)
Time (hours)
Change (fraction of AC)

## Slide 10: Whole blood dilution
Figure 5. Dilution of BOHB-containing whole blood samples in
BOHB-free whole blood. EDTA samples (blue diamonds, ), Heparin sample (green triangles, ). Dashed line is line of identity of expected results.
Interpretation: Dilution in whole blood is not linear.
Predicted BOHB (mmol/L)
BOHB result (mmol/L)

## Slide 11: Sample Addition – low volumes
Figure 6. Addition of low sample volumes. Two heparin whole blood samples tested at progressively lower volumes. Red line shows average of higher volume addition. Note stated minimum is 5 uL.
Interpretation: Addition of sample volumes less than 10 uL may give misleading results. A minimum of 20uL sample volume is recommended.
Sample Volume (uL)
BOHB Result (mmol/L)

## Slide 12: Conclusions
The Abbott Medisense BOHB system is suitable for laboratory use, providing a simple, robust technology for use on EDTA and heparin whole blood samples stored for up to 3 days at 4 degrees or 1 day at room temperature.
Fluoride oxalate whole blood appears to be satisfactory.
All results were within the stated performance criteria.
Care must be taken to ensure sufficient sample volume.
Dilution of high samples is not supported.
Acknowledgement
We thank Abbott Diagnostics for supply of reagents.
References
1. Chiu RKW et al. Evaluation of a new handheld biosensor for point-of-care testing of whole blood beta-hydroxybutyrate concentration. HKMJ 2002;8:172-6
2. Wallace TM et al. The hospital and home use of a 30-second hand-held blood ketone meter: guidelines for clinical practice. Diabetes UK 2001;18:640-5.
3 Guidelines and Recommendations for Laboratory Analysis in the Diagnosis and Management of Diabetes Mellitus. NACB. 2002.