## Slide 1: STANDARDISED PRESENTATION OF INTERFERENCE DATA FOR LIMIT SETTING
BJ Martin, P Graham, M Roser, GRD Jones
Chemical Pathology, St Vincent’s Hospital, Sydney, Australia

## Slide 2: Background
Interference studies are commonly performed for potential interferences from haemolysis, lipaemia and icterus.
The assessment of this data is difficult as there is no standardised criteria for decision making in this area.
In this poster we demonstrate a graphical presentation of interference data to facilitate the process of setting interference limits.

## Slide 3: Methods - Samples
Interference data was generated as follows:
In-vitro haemolysis model:
  Fresh haemolysate at various concentrations from a single donor was added to serum pools.
  Fresh haemolysates from different donors were added to a single serum pool at a single concentration.
Lipaemia Model:
  Intralipid™ was added to serum pools at various concentrations.
Samples were measured on an Olympus AU2700 analyser with user-defined quantitative “H” and “L” indices

## Slide 4: Methods – Olympus Indices
Serum Indices for Haemolysis (H) and Lipaemia (L) were set up on an Olympus AU2700:
  Two additional tests were installed (APP and APP1)
  Wavelengths: APP 570 (primary) / 600 (secondary)
  Wavelengths: APP1 600 (primary) / 700 (secondary)
  Sample volume: 6uL, reagent volume (saline): 125 uL
Calculations
  L = 88.95 * APP1
  H = 3773.12* (APP – (1.25* (APP2))
These quantitative indices were based on Hitachi serum indices and are used in place of the Olympus proprietary semi-quantitative indices.

## Slide 5: Methods - Spreadsheet
Interference data was entered into an Excel spreadsheet application.
Other Inputs
  RCPA-AACB QAP Allowable Limits of Performance (ALP)
  Within-individual Biological Variation (CVwi)
  Instrument / method / setup criteria
  Date of interference data
The aim of the application is to display all the required information for decision making on the same screen or printout.

## Slide 6: Application Display
Graphical displays
  Raw data
  Data expressed as absolute and fractional changes
  Multiple data sets of serial interference addition
    (up to 3 at a time)
  Multiple sets of individual haemolysate preparations
    (up to 18 different haemolysates)
On the same graphs was displayed:
  the RCPA-AACB QAP allowable limits of performance
  The within-person biological variation.
The chart included
  Analytical method, date, catalog number
  Analytical CV

## Slide 7: Application Use
Change analytes for review in 2 or 3 keystrokes
Select data sets by single letter codes
Review current cutoffs and review alternate cutoffs
  Graphically
  Numerically
    (requires selection of data set by single letter code)
Allows direct comparison of data sets according to:
  Instrument
  Date
  Method
  Type

## Slide 9
Test Name
Raw Data
Absolute Change
Relative Change
ALP Color coded
Data set details
Available
Data sets for analyte
Calculated effects
Current & proposed cutoffs
Current Cuttoff
Proposed Cuttoff
Calculation Data

## Slide 11: Conclusions
We have developed a spreadsheet application allowing easy graphical display of multiple data sets for interpretation of interference data.
The data demonstrated has been generated on an Olympus AU2700 using user-defined, quantitative serum indices.
The application allows improved target setting for allowable interference limits.