## Slide 1: Interferences - are some methods better than others?
Graham Jones
Department of Chemical Pathology
St Vincent’s Hospital, Sydney

## Slide 2: Contents
Background
Choosing your instrument
Using your instrument

## Slide 3: Introduction
Our aim: to produce timely, accurate results to allow optimal patient care
Interferences - substances present in a sample, or events affecting a sample, which lead to the production of inaccurate results
Accuracy: results which reflect the result which would have been obtained if the interference had not been present

## Slide 4: Interference Importance
May lead to a clinical error
  Wrong management with bad outcome
Interference-related clinical errors quite rare
  Most clinical errors require several mishaps concurrently
  Many “near misses”
BUT: can cost time, additional testing, reduced doctor confidence

## Slide 5: Error Importance
Erroneous and Non-believable
  eg potassium of 10.0 due to haemolysis or EDTA contamination
  Result: ignore or recollect specimen
Erroneous and Believable
  eg potassium of 5.5 due to haemolysis or EDTA contamination
  result: unnecessarily cease potassium supplements

## Slide 6: Common Interferences
In-vitro haemolysis
Bilirubin
Lipaemia
Drugs
Immunoglobulins
Events (eg delayed separation)
Other (artificial blood)

## Slide 7: Common Interferences
In-vitro haemolysis
Bilirubin
Lipaemia
Drugs
Immunoglobulins
Events (eg delayed separation)
Other (artificial blood)
The visible interferences

## Slide 8: Given factors
We wish to have accurate results
We wish to avoid errors due to interferences
We aim to give out results when they are accurate
We aim to withhold results which are inaccurate
  This implies different cutoff levels for different analytes

## Slide 9: Choosing your instrument

## Slide 10: Assesment of Interferents
Melvin Glick
Clin Chem (1987) 33: 1453-1458
Add known amounts of RBC lysate; Intralipid; bilirubin to normal serum
Standard procedures
Plot percent change in result vs interferent concentration
“Interferographs”

## Slide 11: Interferographs: Glick
100
200
0
Final/original result x 100 (%)
1000
500
0
Haemolysate added (as haemoglobin. mg/dL)
110%
90%
*
Bilirubin
C.Bilirubin
Glucose
* Urea
* Chloride
* Creatinine
GGT

## Slide 12: Glick
Most work performed in 1980s
Work performed using his own blood (reliable supply, but limited quantity)
Limited comprehensive third party data available for current instruments
Data from our own studies
  Haemolysis Interference in Modern Instruments Clin Biochem Revs 2000;21:124
  Icterus Interference in Modern Instruments Clin Biochem Revs 2000;21:124

## Slide 13: Interferogram
Roche Modular <P>
Haemolysis
Haemolysate added to patient samples and concentrations measured

## Slide 14: Instruments

## Slide 15: Comparing Interference Performance: Amylase and Haemolysis
Haemoglobin (mg/dL)
Using RCPA-AACB Allowable Limits of Performance
160
990

## Slide 17: Instrument Comparison
Some Instruments are better than others
  but
All are affected by interferences
Data is NOT transferable between instruments
There is room for improvement by manufacturers

## Slide 18: Effect of Haemolysis - methods
Examples of tests where different instruments show wide variations in response to haemolysis (Data from 2000).

## Slide 19: Method Comparison
Some methods better than others
Suggest choosing methods which are less prone to interference
May require third party supplier

## Slide 20: Using your instrument

## Slide 21: Using Your Instrument
Once the instrument is chosen the fun begins
A protocol must be set which allows appropriate response to samples with interferences
requires detailed knowledge of your method / instrument
Sources:
  Manufacturer
  Literature
  Own studies

## Slide 23
Olympus Cholesterol Reagent and Modular Cholesterol Reagent
Olympus Results
Modular Results

## Slide 25
Olympus Results: 10% at 500
2.5, 4.1 and 6.0 mmol/L
Modular Results
10% at 700
3.5 mmol/L
Response best expressed as absolute (not not percentage)

## Slide 26: Data Sources
Best data is from your own instrument
  No factors
  Full data set
  Perform experiment as needed.
Manufacturer information best when all results available
  Beware of “No Interference” limits (eg 10%)
  Format of limits may not be useful

## Slide 27: How accurate do we need to be?
RCPA-AACB Quality Assurance limits
Change greater than 2 SD of analytical precision
Change related to biological variation
10%
Other fixed percentage or absolute values
A difference that may lead to a change in clinical management - subjective*

## Slide 28: Error Budget
Int. error
Other errors
Total error

## Slide 29: The Accuracy - Utility Balance
More accuracy
More rejections
More recollections
More delays
Unhappier ptns and Drs
Fewer clinical errors
Less accuracy
Fewer rejections
Fewer recollections
Shorter TAT
More clinical errors

## Slide 30: Interference Limits
No easy solution
Take all factors into account
Likely clinical effects is the main parameter
  (personal opinion)
Include pathologist / clinician in decision making

## Slide 31
SydPath Haemolysis Protocol

## Slide 32: Other quality factors
Sample type:
  Serum, heparin plasma, EDTA plasma, fluoride oxalate, Citrate, gel separators.
Sample stability
  As whole blood, as serum/plasma
  At RT, 4 degrees, -20 degrees

## Slide 33
What can I do that will make a difference to your business?
A supplier’s question…

## Slide 34: Suppliers…..
Quality data on interferences, sample types and analyte stability can:
  Reduce recollections
  Reduce unnecessary recollections
  Reduce repetition in multiple laboratories
Head office, literature watch, local data

## Slide 35: Conclusions
Interferences and our response to them are part of providing a quality laboratory service
Choose methods and instruments with low interference
Choose methods where data is available about interferences or generate local data
Implement a policy for responding to interferences