## Slide 1: QUANTIFYING THE EFFECT OF SETTING QUALITY CONTROL STANDARD DEVIATIONS GREATER THAN ACTUAL STANDARD DEVIATIONS ON WESTGARD RULES
Graham Jones
Department of Chemical Pathology,
St Vincent’s Hospital, Sydney, Australia.

## Slide 2: Background
Westgard rules are commonly used to detect changes in assay performances.
The Power of Error Detection (PED) of rules can be determined from available sources1.
These sources assume that the SD set in the QC protocol is the actual SD of the measurement system.
Sometimes the SD in QC charts is set wider than the actual instrument SD (figure 1).
Here I investigate the effect of setting QC limits wider than the actual SDs on the Power of Error Detection.
1. www.westgard.com

## Slide 3: Setting QC Limits - illustration
QC SD = Actual SD
Spread of QC results across range.
5% of results outside +/- 2SD
QC SD = 3 x Actual SD
Same ASD as graph A.
QC results clustered near mean.
No results outside +/- 2SD
Figure 1
+2
0
-2
+2
0
-2
A
B

## Slide 4: Hypothesis
Setting the SD limits in the Levy-Jennings QC chart different to the actual instrument SD limits will change the performance of the QC protocol.
Aim
To investigate the utility of published Power Function Charts to determine the Power of Error Detection when the QCSD is larger than the Actual SD of the method.

## Slide 5: Terminology
ASD - The actual SD of the measurement system
QCSD - The “SD” set in the QC package or drawn on Levy-Jennings chart.
PED: Power of Error Detection - the likelihood that a QC rule will trigger for any given assay drift.
ED90: The error detected with a 90% probability.
13s - 1 result outside 3 SD
12s - 1 result outside 2 SD
22s - 2 results outside 2 SD
41s - 4 consecutive results outside 1 SD
n - the number of QC samples run at the same time

## Slide 6: Methodology
Power Function Charts were developed in Microsoft Excel.
  Random variation was simulated using the random number generator with a normal distribution (n=1000).
  Bias was simulated by adding fixed amounts to the randomly generated numbers.
  Changes in imprecision were modelled by multiplying output from random number generator.
  Rules were modelled for n=2 and n=4.
  Changes in bias and precision were modelled.
  For n= 2, The 22s rule was run within the pair of simultaneous QCs
  For n=4, the 41s rule was run on data from both materials.

## Slide 7: Results and Discussion
Increasing the QCSD relative to the ASD changes the Power Function Charts for detecting bias (n=2, figure 2; n=4, figure 6) and precision (n=2, figure 3).
For changes in bias the PED for various values of QCSD/ASD can be modelled (n=2, figure 4).
For changes in bias the ED90 for various values of QCSD/ASD can be modelled (n=2, fig 5; n=4, fig 6).
  ED90 changes with QCSD/ASD at different rates for different rules and combinations of rules.
  The ED90 can be calculated for various rules from the data presented (see figures 5 and 6 for data).

## Slide 8
PED
Power Function Chart (bias, n=2)
Effect of increasing QCSD relative to ASD on Power of Error Detection (PED) for various changes in bias. Rules: 13s/22s/R4s. n=2. ED90 is used for later calculations (fig 5) Conclusion: PED for bias falls as QCSD/ASD increases.
Figure 2
Shift (multiples of SD)
QCSD/ASD ED90

## Slide 9: Power Function Charts (precision, n=2)
PED
Precision (multiples of SD)
Effect of increasing QCSD / ASD on the Power of Error Detection (PED) for changes in assay precision. Rules: 13s/22s/R4s. N=2. Note: PED never reaches 90%. Conclusion: PED for precision falls as QCSD/ASD increases.
QCSD/ASD
Figure 3

## Slide 10: Increasing QCSD / ASD: Effect on PED
QCSD / ASD
PED
Figure 4
Effect of increasing QCSD/ASD on the Power of Error Detection for various shifts in assay bias. Rules: 13s/22s/R4s. n=2.
Conclusion: PED falls as QCSD/ASD increases.
Shift
(multiples of ASD)

## Slide 11: Increasing QCSD / ASD: Effect on ED90
QCSD / ASD
ED90
Figure 5
Effect of increasing QCSD/ASD on the Error Detection with 90% probability (ED90)for individual rules and combinations of rules. n=2.
Conclusion: ED90 increases linearly with increasing QCSD/ASD.
13s/22s/R4s ED90=2.3 x QCSD/ASD + 0.9
22s ED90=2.0 x QCSD/ASD + 1.6
13s ED90=3.0 x QCSD/ASD + 0.4
ED90 Prediction Equations

## Slide 12
PED
Power Function Charts (bias, n=4)
A. Effect of increasing QCSD / ASD on PED. N=4. Rules: 13s/22s/R4s/41s. n=4.
B. ED90 for various rules. n=4.
Figure 6
Shift (multiples of SD)
QCSD/ASD ED90
13s/22s/41s/R4s ED90=1.5 x QCSD/ASD + 0.8
13s ED90 = 3.0 x QCSD/ASD - 0.1
22s ED90 = 2.0 x QCSD/ASD + 0.6
41s ED90 = 1.0 x QCSD/ASD + 2
ED90
A
B
QCSD / ASD
Conclusion: Data similar for n=4.
ED90 Prediction Equations

## Slide 13: Conclusions
Setting the QCSD wider than the ASD in QC protocols affects the performance of Westgard rules.
If QCSD does not equal to ASD the published Power Function Charts do not accurately represent the power of error detection.
Laboratories wishing to determine the power of error detection using published data must either set QCSD=ASD; apply corrections as outlined in this poster; or develop alternate methods to determine their own error detection power.